## KingstonL
Subject |  | 001-WN | 002-AS | 003-VA
Age |  | 39 | 33 | 30
Height |  | 169 | 185 | 170
Weight |  | 55 | 85 | 80
BMI
Gender |  | M | M | M
QNFT AccuFIT9000, 2020-04-22, AS, WN, VA
Mask |  | KingstonL S | KingstonL M | KingstonL M
Filter |  | Vacuum 2 Layer | Vacuum 2 Layer | Vacuum 2 Layer
Fit Factor
Normal Breathing |  | 6.76 | 10.48 | 12.47
Deep Breathing |  | 8.76 | 9.66 | 17.26
Side to Side |  | 8.07 | 10.44 | 10.01
Nodding Head |  | 8.24 | 11.51 | 11.06
Talking Out Loud |  | 7.32 | 8 | 10.2
Bending Over |  | 7.78 | 13.61 | 7.03
Normal Breathing |  | 8.81 | 16.12 | 5.95
Grimace not-done |  | Null | Null | Null
Overall Fit Factor |  | 7.9 | 10.9 | 9.08
Average Fit Factor |  | 7.9629 | 11.4029 | 10.5686
Ease of Breathing | 5 is very easy |  |  | 3
Comfort | 5 is very comfortable |  |  | 2
Notes: |  | Stretching Elastic Band to fit - sensed a leak around his cheeks. Mask digging into sides of his nose |  | Straps cut and tied. Still had to hold tightly for best possible fit
"Hi again all,
We have done quantitative fit testing the 3 models you’ve sent us (Kingston, KingstonL, SnowCone), on 3 different people. The Kingston models were both tested with 2 layers of Vacuum MERV 15 filter, the SnowCone was tested with 1 layer, as we were unable to get 2 layers in. For QNFT testing according to CSA standards, the fit factor must be at least 100. The results (overall fit factors) were as follows:
Kingston: 8.43, 11.76, 8.42
KingstonL: 7.90, 10.90, 9.08
SnowCone: 4.57, 5.18, 3.01
We have also scored the masks on a scale of 1-5 for both ease of breathing, and comfort.
We will upload the spreadsheet to GitHub for you to look at - it includes the results for each of the 7 separate tests that make up the QNFT testing.
Sorry to disappoint. These are the same kind of results we were getting with our masks that we tested last week - despite feeling like a decent fit, the machine doesn’t agree.
Andrew"

## Kingston
Subject |  | 001-WN | 002-AS | 003-VA
Age |  | 39 | 33 | 30
Height |  | 169 | 185 | 170
Weight |  | 55 | 85 | 80
BMI
Gender |  | M | M | M
QNFT AccuFIT9000, 2020-04-22, AS, WN, VA
Mask |  | Kingston | Kingston | Kingston M
Filter |  | Vacuum 2 Layer | Vacuum 2 Layer | Vacuum 2 Layer
Fit Factor
Normal Breathing |  | 12.74 | 12.19 | 9.49
Deep Breathing |  | 6.98 | 7.95 | 8.98
Side to Side |  | 8.68 | 14.27 | 7.37
Nodding Head |  | 6.2 | 12.99 | 7.5
Talking Out Loud |  | 7.26 | 10 | 9.31
Bending Over |  | 9.51 | 12.86 | 6.85
Normal Breathing |  | 11.23 | 15.98 | 10.82
Grimace not-done |  | Null | Null | Null
Overall Fit Factor |  | 8.43 | 11.76 | 8.42
Average |  | 8.9429 | 12.32 | 8.6171
Ease of Breathing | 5 is very easy | 4 | 2 | 3
Comfort | 5 is very comfortable | 2 | 3 | 4
Comments |  | Leaking around nose, marks over nose
Comments |  | Digging into nose. Can feel leaks at the sides of cheeks |  | Holding straps at the back

## SnowCone
Subject |  | 001-WN | 002-AS | 003-VA
Age |  | 39 | 33 | 30
Height |  | 169 | 185 | 170
Weight |  | 55 | 85 | 80
BMI
Gender |  | M | M | M
QNFT AccuFIT9000, 2020-04-22, AS, WN, VA
Mask |  | SnowCone S | SnowCone S | SnowCone S
Filter |  | Vacuum 1 Layer | Vacuum 1 Layer | Vacuum 1 Layer
Fit Factor
Normal Breathing |  | 4.3 | 5.73 | 2.78
Deep Breathing |  | 4.56 | 4.73 | 2.82
Side to Side |  | 4.52 | 5.12 | 2.81
Nodding Head |  | 4.64 | 4.78 | 2.93
Talking Out Loud |  | 4.17 | 6.47 | 4.57
Bending Over |  | 4.94 | 4.49 | 2.89
Normal Breathing |  | 5.01 | 5.46 | 2.81
Grimace not-done |  | Null | Null | Null
Overall Fit Factor |  | 4.57 | 5.18 | 3.01
Average |  | 4.5914 | 5.2543 | 3.0871
Ease of Breathing | 5 is very easy | 5 | 4 | 4
Comfort | 5 is very comfortable | 2 | 2 | 1
Comments |  | Digging into nose. Very uncomfortable. Leads around sides of cheeks
Comments |  | Digs into nose. Not much of leak with hand test on expiration.
Note on C12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: MERV 15