## Manitoba DS30 DS30
AccuFIT-9000 testing, CSA QNFT protocol
Location: TGH Anesthesia
Date: 2020-04-25
Testers: AS, WN
Subject
Demographic | 001-WN | 004-WN | 001-AS | 001-VA | 002-WN | 003-WN
Age | 39 | 39 | 33 | 30
Height | 169 | 169 | 185 | 170
Weight | 55 | 55 | 85 | 80
BMI
Gender | M | M | M | M
| Nasal Rim Tear, No Touch |  |  |  | Will's repeat testing trying to get perfect seal, Sample Tube Falls
Mask | Manitoba DS30/DS10 | Manitoba DS30/DS10 | Manitoba DS30/DS30 |  | Manitoba DS30/DS10
Filter | InterSurgical HME |  | InterSurgical HME |  | InterSurgical HME
Mask Colour | Green | White |  |  | Green | Green
Run | Fit Factor
Normal Breathing | 42.17 | 195.59 |  |  | 43350.85 | 224.73
Deep Breathing | 35.79 | 118.08 |  |  | 14444.03 | 212.57
Side to Side | 36.53 | 92.77 |  |  | 3.02 | 50.44
Nodding Head | 63.39 | 127.24 |  |  | 193.26 | 68.23
Talking Out Loud | 57.8 | 83.13 |  |  | 2133.43 | N/A
Bending Over | 54.96 | 92.06 |  |  | 1570.36 | N/A
Normal Breathing | 27.82 | 104.19 |  |  | 1036.23 | N/A
Grimace not-done | N/A | N/A |  |  | N/A | N/A
Overall Fit Factor | 42.11 | 108.18 |  |  | N/A | N/A
Average | 45.4943 | 116.1514 |  |  | N/A | N/A
Ease of Breathing | 4 | 4
Comfort | 2 | 2
Comments |  | Best fit so far - no nose pinching |  |  | Pinching nose closed | Pinching nose closed
Comments | DS10 Straps, narrow, tight
Comments | Leak around nasal bridge due to tear |  |  |  | Sampling tubing fell off during side to side and nodding
Note on A31: Microsoft Office User: 5 is very easy
Note on A32: Microsoft Office User: 5 is very comfortable